## Flow
| Flow 1 |  |  | Summary:
| *This is Current implemted Version. It Take 22sec to complete a Sweep for 100ms Sweepdelay |  |  | 1. The Board need time to Discharge whevever there is range change
| *1 Rank Range search from Largers to smallest range (20mA to 2nA) |  |  | 2. Discharge time may differ base on the Gap Between Range and Actual Measurement
|  | Total TestTime : ~22sec (0 to3V, 0.1v Stepdelay) |  | 3.Communication time within connector and device take about 50ms for each command set
Improvement:
1.Autorange to be done at firmware level to have better test time control.
Autorange at software level is not a good approach as the Communication time between the board and connector will be influnce by the PC speed or interupt by other pc's process.
2.Improve on the board Discharge time, faster discharge time determine how fast the Autorange process
3.Improve board communication response time
| Flow 2 -Improved Test time |  | Do Not Applicable to current application.
According to JB suggestion:
a) | Set only one time bias Voltage , then only search for suitable current range |  | *This Flow do not do Range Search, Range is apply base on the Saved Range
b) | 1 Rank Range Search start from Lowest to Largest Range (2nA - 20mA)
c) | No SweepDelay During Range Search, After found Range, then only followy by Sweep delay and fetch new Measurement again
d) | Stay at previous Current Range for next voltage step
|  | Start Test Flow |  |  | Replica Test Flow
|  | Total TestTime : ~17sec (0 to3V, 0.1v Stepdelay) |  |  | Total TestTime : ~11sec (0 to3V, 0.1v Stepdelay)
Flow 3
Changes from Flow 2 According to 1st July Meeting
a) | Save the Sweepdelay Time with minus Process time
If SweepDelay time more than Process time , then wait for remaining sweepdelay and do new measurement again
if sweepDelay time less than Process time , then the last measurement result during autorange search will be use for final result
Start
Set BiasVoltage =>03
Set Current Range =>0A
StartTest=>05
Fetch Measurement=>08
Correct Current Range?
[Search frm Lowest Range]
Voltage > = Stop Voltage?
End
Voltage =+Step Voltage
Yes
No
YES
Set to Next 1
Higher Range
No
50ms to wait Response
50ms to wait Response
50ms to wait Response
50ms to wait Response
StepDelay
StartTest=>05
Fetch Measurement=>08
50ms to wait Response
50ms to wait Response
Start
Set BiasVoltage =>03
Set Current Range =>0A
StartTest=>05
Fetch Measurement=>08
Voltage > = Stop Voltage?
End
Voltage =+Step Voltage
Yes
No
50ms to wait Response
50ms to wait Response
50ms to wait Response
50ms to wait Response
StepDelay
Start
Set BiasVoltage =>03
Set Current Range =>0A
StartTest=>05
Fetch Measurement=>08
Correct Current Range?
[Search frm Lowest Range]
Voltage > = Stop Voltage?
End
Voltage =+Step Voltage
Yes
No
YES
Set to Next 1
Higher Range
No
50ms to wait Response
50ms to wait Response
50ms to wait Response
50ms to wait Response
SweepDelay =SweepDelay -Process time
StartTest=>05
Fetch Measurement=>08
50ms to wait Response
50ms to wait Response
Is SweepDelay >ProcessTime?
YES
No
50ms to wait Response
50ms to wait Response
50ms to wait Response
50ms to wait Response
Start
Set Current Range =>0A
Set BiasVoltage =>03
StartTest=>05
Fetch Measurement=>08
Correct Current Range?
[Search from Highest Range]
Voltage > = Stop Voltage?
End
Voltage =+Step Voltage
Yes
No
YES
Set to Next Lower Range
No
sweepDelay

## 1RangeSearch-Flow2
[8 empty] | AD Full Scale Voltage (V) : | 3.3
[8 empty] | AD Reference Voltage Offset (V) : | 1.5
[8 empty] | AD Resolution (bits) : | 24
[8 empty] | Computational Constant : | 3813003.6364
[8 empty] | Measurement Range | Multiplier | [6 empty] | Note:
[8 empty] | 1 | 2e-09 |  | (2nA) |  |  |  |  | Autorange in 1 Range up from low to high Range using Flow 2
[8 empty] | 2 | 2e-08 |  | (20nA) |  |  |  |  | sweepdelay=100ms
[8 empty] | 3 | 2e-07 |  | (200nA)
[8 empty] | 4 | 2e-06 |  | (2uA)
[8 empty] | 5 | 2e-05 |  | (20uA)
[8 empty] | 6 | 0.0002 |  | (200uA)
[8 empty] | 7 | 0.002 |  | (2mA)
[8 empty] | 8 | 0.02 |  | (20mA)
Test Results Using Diode and 150 ohm Series Resistor, Forward Bias (Dummy Specimen) Date : 11 May 2022
Hexadecimal Data From Device (Fetch Measurement) | [11 empty] | Bias Voltage (mV) | Measuement Range (A) | Unsigned Raw AD Value | Signed Raw AD Value | Resultant Bias Current (A)
1 | 8 | 8 | 0 | 0 | EE | E7 | DF | 1 | 93 | E9 |  | 0 | 2e-09 | 15656927 | -1120289 | 5.8761e-10
1 | 8 | 8 | 0 | 0 | FC | 43 | 87 | 1 | EC | B2 |  | 0 | 2e-09 | 16532359 | -244857 | 1.2843e-10
1 | 8 | 8 | 64 | 0 | D4 | CE | 26 | 1 | 4 | 8D |  | 100 | 2e-09 | 13946406 | -2830810 | 1.4848e-09
1 | 8 | 8 | 64 | 0 | BE | DD | EA | 1 | BD | 90 |  | 100 | 2e-09 | 12508650 | -4268566 | 2.239e-09
1 | 8 | 8 | C8 | 0 | BE | DD | F0 | 1 | AE | 9C |  | 200 | 2e-09 | 12508656 | -4268560 | 2.2389e-09
1 | 8 | 8 | C8 | 0 | BE | DD | F0 | 2 | EE | 9D |  | 200 | 2e-08 | 12508656 | -4268560 | 2.2389e-08
1 | 8 | 8 | C8 | 0 | C2 | F6 | B9 | 3 | 71 | 55 |  | 200 | 2e-07 | 12777145 | -4000071 | 2.0981e-07
1 | 8 | 8 | C8 | 0 | C0 | 46 | F5 | 4 | 4 | 8 |  | 200 | 2e-06 | 12601077 | -4176139 | 2.1905e-06
1 | 8 | 8 | C8 | 0 | ED | A9 | 56 | 5 | 85 | A1 |  | 200 | 2e-05 | 15575382 | -1201834 | 6.3039e-06
1 | 8 | 8 | C8 | 0 | FF | 38 | 74 | 5 | C9 | 94 |  | 200 | 2e-05 | 16726132 | -51084 | 2.6795e-07
1 | 8 | 8 | 2C | 1 | FD | 9F | 1 | 5 | 75 | 7B |  | 300 | 2e-05 | 16621313 | -155903 | 8.1774e-07
1 | 8 | 8 | 2C | 1 | F0 | E3 | 60 | 4 | 5E | 5F |  | 300 | 2e-06 | 15786848 | -990368 | 5.1947e-07
1 | 8 | 8 | 2C | 1 | C9 | F6 | 87 | 4 | 8 | F7 |  | 300 | 2e-06 | 13235847 | -3541369 | 1.8575e-06
1 | 8 | 8 | 90 | 1 | C0 | F5 | 9B | 4 | E9 | 57 |  | 400 | 2e-06 | 12645787 | -4131429 | 2.167e-06
1 | 8 | 8 | 90 | 1 | C0 | 2F | BD | 5 | 13 | 0C |  | 400 | 2e-05 | 12595133 | -4182083 | 2.1936e-05
1 | 8 | 8 | 90 | 1 | D3 | 11 | 74 | 6 | 60 | 15 |  | 400 | 0.0002 | 13832564 | -2944652 | 0.00015445
1 | 8 | 8 | 90 | 1 | FB | 81 | 6F | 6 | 63 | 68 |  | 400 | 0.0002 | 16482671 | -294545 | 1.545e-05
1 | 8 | 8 | F4 | 1 | E4 | 55 | 5F | 6 | 38 | A0 |  | 500 | 0.0002 | 14964063 | -1813153 | 9.5104e-05
1 | 8 | 8 | F4 | 1 | E3 | 3E | D9 | 6 | 2A | 68 |  | 500 | 0.0002 | 14892761 | -1884455 | 9.8844e-05
1 | 8 | 8 | 58 | 2 | D3 | 1A | 26 | 6 | 78 | FF |  | 600 | 0.0002 | 13834790 | -2942426 | 0.00015434
1 | 8 | 8 | 58 | 2 | C0 | D3 | 7F | 6 | 96 | 15 |  | 600 | 0.0002 | 12637055 | -4140161 | 0.00021716
1 | 8 | 8 | BC | 2 | C0 | D2 | 38 | 6 | E3 | C1 |  | 700 | 0.0002 | 12636728 | -4140488 | 0.00021718
1 | 8 | 8 | BC | 2 | C0 | 3F | DA | 7 | FA | 94 |  | 700 | 0.002 | 12599258 | -4177958 | 0.0021914
1 | 8 | 8 | BC | 2 | E6 | 28 | 60 | 8 | 72 | BC |  | 700 | 0.02 | 15083616 | -1693600 | 0.0088833 | Not Fully Discharged, when range change from 2mA to 20mA
1 | 8 | 8 | BC | 2 | FD | BE | 92 | 8 | D1 | 14 |  | 700 | 0.02 | 16629394 | -147822 | 0.00077536
1 | 8 | 8 | 20 | 3 | FD | 36 | 6A | 8 | 32 | A2 |  | 800 | 0.02 | 16594538 | -182678 | 0.00095818
1 | 8 | 8 | 20 | 3 | F5 | BF | 49 | 7 | B8 | 1C |  | 800 | 0.002 | 16105289 | -671927 | 0.00035244
1 | 8 | 8 | 20 | 3 | DC | 89 | B7 | 7 | 10 | 2E |  | 800 | 0.002 | 14453175 | -2324041 | 0.001219
1 | 8 | 8 | 84 | 3 | D5 | E1 | BC | 7 | 8C | 7A |  | 900 | 0.002 | 14016956 | -2760260 | 0.0014478
1 | 8 | 8 | 84 | 3 | CA | 2E | 12 | 7 | C7 | F1 |  | 900 | 0.002 | 13250066 | -3527150 | 0.0018501
1 | 8 | 8 | E8 | 3 | C9 | 31 | A3 | 7 | 8B | 4F |  | 1000 | 0.002 | 13185443 | -3591773 | 0.001884
1 | 8 | 8 | E8 | 3 | BE | D1 | 10 | 7 | E4 | 3D |  | 1000 | 0.002 | 12505360 | -4271856 | 0.0022407
1 | 8 | 8 | 4C | 4 | BE | D2 | 2 | 7 | B4 | 79 |  | 1100 | 0.002 | 12505602 | -4271614 | 0.0022406
1 | 8 | 8 | 4C | 4 | C4 | E8 | 0E | 8 | C8 | 68 |  | 1100 | 0.02 | 12904462 | -3872754 | 0.020313 | OverShoot even at Highest Range,Not yet Discharging
1 | 8 | 8 | 4C | 4 | F7 | 11 | 92 | 8 | 7E | 1D |  | 1100 | 0.02 | 16191890 | -585326 | 0.0030702
1 | 8 | 8 | B0 | 4 | F5 | 42 | 2F | 8 | EA | D8 |  | 1200 | 0.02 | 16073263 | -703953 | 0.0036924
1 | 8 | 8 | B0 | 4 | F5 | 42 | 0C | 8 | F3 | E8 |  | 1200 | 0.02 | 16073228 | -703988 | 0.0036926
1 | 8 | 8 | 14 | 5 | F3 | F4 | BC | 8 | 43 | A2 |  | 1300 | 0.02 | 15987900 | -789316 | 0.0041401
1 | 8 | 8 | 14 | 5 | F3 | 6F | F1 | 8 | 7 | 1D |  | 1300 | 0.02 | 15953905 | -823311 | 0.0043184
1 | 8 | 8 | 78 | 5 | F1 | 93 | 48 | 8 | BC | 69 |  | 1400 | 0.02 | 15831880 | -945336 | 0.0049585
1 | 8 | 8 | 78 | 5 | F1 | 93 | 34 | 8 | 9C | A9 |  | 1400 | 0.02 | 15831860 | -945356 | 0.0049586
1 | 8 | 8 | DC | 5 | F0 | 21 | 3B | 8 | 21 | A6 |  | 1500 | 0.02 | 15737147 | -1040069 | 0.0054554
1 | 8 | 8 | DC | 5 | EF | AD | C5 | 8 | A7 | F9 |  | 1500 | 0.02 | 15707589 | -1069627 | 0.0056104
1 | 8 | 8 | 40 | 6 | EF | 1C | 4D | 8 | C8 | 42 |  | 1600 | 0.02 | 15670349 | -1106867 | 0.0058057
1 | 8 | 8 | 40 | 6 | ED | C2 | 89 | 8 | FB | 0 |  | 1600 | 0.02 | 15581833 | -1195383 | 0.00627
1 | 8 | 8 | A4 | 6 | EC | 2D | 14 | 8 | B4 | 7D |  | 1700 | 0.02 | 15478036 | -1299180 | 0.0068145
1 | 8 | 8 | A4 | 6 | EB | D6 | 82 | 8 | AB | 58 |  | 1700 | 0.02 | 15455874 | -1321342 | 0.0069307
1 | 8 | 8 | 8 | 7 | EB | 31 | 88 | 8 | 38 | 63 |  | 1800 | 0.02 | 15413640 | -1363576 | 0.0071522
1 | 8 | 8 | 8 | 7 | E9 | EF | 3D | 8 | 2F | 71 |  | 1800 | 0.02 | 15331133 | -1446083 | 0.007585
1 | 8 | 8 | 6C | 7 | E8 | 36 | BC | 8 | 97 | 2 |  | 1900 | 0.02 | 15218364 | -1558852 | 0.0081765
1 | 8 | 8 | 6C | 7 | E7 | FA | 5 | 8 | 27 | B9 |  | 1900 | 0.02 | 15202821 | -1574395 | 0.008258
1 | 8 | 8 | D0 | 7 | E7 | 38 | 11 | 8 | 93 | B9 |  | 2000 | 0.02 | 15153169 | -1624047 | 0.0085185
1 | 8 | 8 | D0 | 7 | E6 | 0D | 46 | 8 | BC | 7B |  | 2000 | 0.02 | 15076678 | -1700538 | 0.0089197
1 | 8 | 8 | 34 | 8 | E4 | 4A | EE | 8 | 31 | F3 |  | 2100 | 0.02 | 14961390 | -1815826 | 0.0095244
1 | 8 | 8 | 34 | 8 | E4 | 21 | 7 | 8 | 0F | BF |  | 2100 | 0.02 | 14950663 | -1826553 | 0.0095807
1 | 8 | 8 | 98 | 8 | E3 | 4E | 7C | 8 | 4 | 4A |  | 2200 | 0.02 | 14896764 | -1880452 | 0.0098634
1 | 8 | 8 | 98 | 8 | E2 | 33 | 43 | 8 | 84 | 5E |  | 2200 | 0.02 | 14824259 | -1952957 | 0.010244
1 | 8 | 8 | FC | 8 | E0 | 57 | 4A | 8 | CA | 4D |  | 2300 | 0.02 | 14702410 | -2074806 | 0.010883
1 | 8 | 8 | FC | 8 | E0 | 3E | AB | 8 | 52 | 1 |  | 2300 | 0.02 | 14696107 | -2081109 | 0.010916
1 | 8 | 8 | 60 | 9 | DF | 50 | 74 | 8 | 47 | A4 |  | 2400 | 0.02 | 14635124 | -2142092 | 0.011236
1 | 8 | 8 | 60 | 9 | DE | 48 | 61 | 8 | C8 | CF |  | 2400 | 0.02 | 14567521 | -2209695 | 0.01159
1 | 8 | 8 | C4 | 9 | DC | 4B | D9 | 8 | 52 | 53 |  | 2500 | 0.02 | 14437337 | -2339879 | 0.012273
1 | 8 | 8 | C4 | 9 | DC | 4B | 31 | 8 | 1C | 53 |  | 2500 | 0.02 | 14437169 | -2340047 | 0.012274
1 | 8 | 8 | 28 | 0A | DB | 43 | A5 | 8 | A0 | 89 |  | 2600 | 0.02 | 14369701 | -2407515 | 0.012628
1 | 8 | 8 | 28 | 0A | DA | 5C | 2B | 8 | F5 | 13 |  | 2600 | 0.02 | 14310443 | -2466773 | 0.012939
1 | 8 | 8 | 8C | 0A | D8 | 65 | 1A | 8 | 28 | 12 |  | 2700 | 0.02 | 14181658 | -2595558 | 0.013614
1 | 8 | 8 | 8C | 0A | D8 | 64 | C5 | 8 | 21 | E2 |  | 2700 | 0.02 | 14181573 | -2595643 | 0.013615
1 | 8 | 8 | F0 | 0A | D7 | 3C | 4D | 8 | CE | D9 |  | 2800 | 0.02 | 14105677 | -2671539 | 0.014013
1 | 8 | 8 | F0 | 0A | D6 | 6D | DE | 8 | F2 | 4 |  | 2800 | 0.02 | 14052830 | -2724386 | 0.01429
1 | 8 | 8 | 54 | 0B | D4 | 78 | C9 | 8 | C9 | 6C |  | 2900 | 0.02 | 13924553 | -2852663 | 0.014963
1 | 8 | 8 | 54 | 0B | D4 | 77 | EE | 8 | E2 | 9F |  | 2900 | 0.02 | 13924334 | -2852882 | 0.014964
1 | 8 | 8 | B8 | 0B | D3 | 3E | 5C | 8 | 50 | 31 |  | 3000 | 0.02 | 13844060 | -2933156 | 0.015385
1 | 8 | 8 | B8 | 0B | D2 | 89 | 3F | 8 | C9 | 1B |  | 3000 | 0.02 | 13797695 | -2979521 | 0.015628
Test Results Using Diode and 150 ohm Series Resistor, Forward Bias (Dummy Specimen) Date : 11 May 2022
Hexadecimal Data From Device (Fetch Measurement) | [11 empty] | Bias Voltage (mV) | Measuement Range (A) | Unsigned Raw AD Value | Signed Raw AD Value | Resultant Bias Current (A)
1 | 8 | 8 | 0 | 0 | FC | 43 | 87 | 1 | EC | B2 |  | 0 | 2e-09 | 16532359 | -244857 | 1.2843e-10
1 | 8 | 8 | 64 | 0 | BE | DD | EA | 1 | BD | 90 |  | 100 | 2e-09 | 12508650 | -4268566 | 2.239e-09
1 | 8 | 8 | C8 | 0 | FF | 38 | 74 | 5 | C9 | 94 |  | 200 | 2e-05 | 16726132 | -51084 | 2.6795e-07
1 | 8 | 8 | 2C | 1 | C9 | F6 | 87 | 4 | 8 | F7 |  | 300 | 2e-06 | 13235847 | -3541369 | 1.8575e-06
1 | 8 | 8 | 90 | 1 | FB | 81 | 6F | 6 | 63 | 68 |  | 400 | 0.0002 | 16482671 | -294545 | 1.545e-05
1 | 8 | 8 | F4 | 1 | E3 | 3E | D9 | 6 | 2A | 68 |  | 500 | 0.0002 | 14892761 | -1884455 | 9.8844e-05
1 | 8 | 8 | 58 | 2 | C0 | D3 | 7F | 6 | 96 | 15 |  | 600 | 0.0002 | 12637055 | -4140161 | 0.00021716
1 | 8 | 8 | BC | 2 | FD | BE | 92 | 8 | D1 | 14 |  | 700 | 0.02 | 16629394 | -147822 | 0.00077536
1 | 8 | 8 | 20 | 3 | DC | 89 | B7 | 7 | 10 | 2E |  | 800 | 0.002 | 14453175 | -2324041 | 0.001219
1 | 8 | 8 | 84 | 3 | CA | 2E | 12 | 7 | C7 | F1 |  | 900 | 0.002 | 13250066 | -3527150 | 0.0018501
1 | 8 | 8 | E8 | 3 | BE | D1 | 10 | 7 | E4 | 3D |  | 1000 | 0.002 | 12505360 | -4271856 | 0.0022407
1 | 8 | 8 | 4C | 4 | F7 | 11 | 92 | 8 | 7E | 1D |  | 1100 | 0.02 | 16191890 | -585326 | 0.0030702
1 | 8 | 8 | B0 | 4 | F5 | 42 | 0C | 8 | F3 | E8 |  | 1200 | 0.02 | 16073228 | -703988 | 0.0036926
1 | 8 | 8 | 14 | 5 | F3 | 6F | F1 | 8 | 7 | 1D |  | 1300 | 0.02 | 15953905 | -823311 | 0.0043184
1 | 8 | 8 | 78 | 5 | F1 | 93 | 34 | 8 | 9C | A9 |  | 1400 | 0.02 | 15831860 | -945356 | 0.0049586
1 | 8 | 8 | DC | 5 | EF | AD | C5 | 8 | A7 | F9 |  | 1500 | 0.02 | 15707589 | -1069627 | 0.0056104
1 | 8 | 8 | 40 | 6 | ED | C2 | 89 | 8 | FB | 0 |  | 1600 | 0.02 | 15581833 | -1195383 | 0.00627
1 | 8 | 8 | A4 | 6 | EB | D6 | 82 | 8 | AB | 58 |  | 1700 | 0.02 | 15455874 | -1321342 | 0.0069307
1 | 8 | 8 | 8 | 7 | E9 | EF | 3D | 8 | 2F | 71 |  | 1800 | 0.02 | 15331133 | -1446083 | 0.007585
1 | 8 | 8 | 6C | 7 | E7 | FA | 5 | 8 | 27 | B9 |  | 1900 | 0.02 | 15202821 | -1574395 | 0.008258
1 | 8 | 8 | D0 | 7 | E6 | 0D | 46 | 8 | BC | 7B |  | 2000 | 0.02 | 15076678 | -1700538 | 0.0089197
1 | 8 | 8 | 34 | 8 | E4 | 21 | 7 | 8 | 0F | BF |  | 2100 | 0.02 | 14950663 | -1826553 | 0.0095807
1 | 8 | 8 | 98 | 8 | E2 | 33 | 43 | 8 | 84 | 5E |  | 2200 | 0.02 | 14824259 | -1952957 | 0.010244
1 | 8 | 8 | FC | 8 | E0 | 3E | AB | 8 | 52 | 1 |  | 2300 | 0.02 | 14696107 | -2081109 | 0.010916
1 | 8 | 8 | 60 | 9 | DE | 48 | 61 | 8 | C8 | CF |  | 2400 | 0.02 | 14567521 | -2209695 | 0.01159
1 | 8 | 8 | C4 | 9 | DC | 4B | 31 | 8 | 1C | 53 |  | 2500 | 0.02 | 14437169 | -2340047 | 0.012274
1 | 8 | 8 | 28 | 0A | DA | 5C | 2B | 8 | F5 | 13 |  | 2600 | 0.02 | 14310443 | -2466773 | 0.012939
1 | 8 | 8 | 8C | 0A | D8 | 64 | C5 | 8 | 21 | E2 |  | 2700 | 0.02 | 14181573 | -2595643 | 0.013615
1 | 8 | 8 | F0 | 0A | D6 | 6D | DE | 8 | F2 | 4 |  | 2800 | 0.02 | 14052830 | -2724386 | 0.01429
1 | 8 | 8 | 54 | 0B | D4 | 77 | EE | 8 | E2 | 9F |  | 2900 | 0.02 | 13924334 | -2852882 | 0.014964
1 | 8 | 8 | B8 | 0B | D2 | 89 | 3F | 8 | C9 | 1B |  | 3000 | 0.02 | 13797695 | -2979521 | 0.015628
Problem 1 : there is time required to discharge to actual current whenever there is current range change. The larger the Gap between actual and Range might need long time to discharge, this issue will also cause inaccurate for the Autorange to pick correct Range to apply.

## 1RangeSearch-Flow3
[8 empty] | AD Full Scale Voltage (V) : | 3.3
[8 empty] | AD Reference Voltage Offset (V) : | 1.5
[8 empty] | AD Resolution (bits) : | 24
[8 empty] | Computational Constant : | 3813003.6364 | [6 empty] | Note:
Autorange in 1 Range up from low to high Range using Flow3
[8 empty] | Measurement Range | Multiplier | [6 empty] | sweepdelay=100ms
[8 empty] | 1 | 2e-09 |  | (2nA)
[8 empty] | 2 | 2e-08 |  | (20nA)
[8 empty] | 3 | 2e-07 |  | (200nA)
[8 empty] | 4 | 2e-06 |  | (2uA)
[8 empty] | 5 | 2e-05 |  | (20uA)
[8 empty] | 6 | 0.0002 |  | (200uA)
[8 empty] | 7 | 0.002 |  | (2mA)
[8 empty] | 8 | 0.02 |  | (20mA)
Test Results Using Diode and 150 ohm Series Resistor, Forward Bias (Dummy Specimen) Date : 11 May 2022
Hexadecimal Data From Device (Fetch Measurement) | [11 empty] | Bias Voltage (mV) | Measuement Range (A) | Unsigned Raw AD Value | Signed Raw AD Value | Resultant Bias Current (A)
1 | 8 | 8 | 0 | 0 | 0 | 4 | BD | 1 | 7E | 57 |  | 0 | 2e-09 | 1213 | 1213 | -6.3624e-13
1 | 8 | 8 | 64 | 0 | E4 | 1 | 9E | 1 | 49 | B2 |  | 100 | 2e-09 | 14942622 | -1834594 | 9.6228e-10
1 | 8 | 8 | C8 | 0 | BE | DD | D3 | 1 | B7 | AC |  | 200 | 2e-09 | 12508627 | -4268589 | 2.239e-09
1 | 8 | 8 | C8 | 0 | BE | DD | D1 | 2 | F6 | CD |  | 200 | 2e-08 | 12508625 | -4268591 | 2.239e-08
1 | 8 | 8 | C8 | 0 | BE | DD | CC | 3 | 3E | 5D |  | 200 | 2e-07 | 12508620 | -4268596 | 2.239e-07
1 | 8 | 8 | C8 | 0 | C0 | CB | 8 | 4 | D4 | B3 |  | 200 | 2e-06 | 12634888 | -4142328 | 2.1727e-06
1 | 8 | 8 | C8 | 0 | F4 | DB | 38 | 5 | 0E | 86 |  | 200 | 2e-05 | 16046904 | -730312 | 3.8306e-06
1 | 8 | 8 | 2C | 1 | F9 | 3B | D8 | 5 | 6E | 38 |  | 300 | 2e-05 | 16333784 | -443432 | 2.3259e-06
1 | 8 | 8 | 90 | 1 | E8 | AD | E5 | 5 | 81 | 44 |  | 400 | 2e-05 | 15248869 | -1528347 | 8.0165e-06
1 | 8 | 8 | F4 | 1 | C4 | A8 | 83 | 5 | BA | 51 |  | 500 | 2e-05 | 12888195 | -3889021 | 2.0399e-05
1 | 8 | 8 | F4 | 1 | BE | DD | CB | 6 | C4 | 92 |  | 500 | 0.0002 | 12508619 | -4268597 | 0.0002239
1 | 8 | 8 | F4 | 1 | E0 | 78 | 25 | 7 | 4B | 39 |  | 500 | 0.002 | 14710821 | -2066395 | 0.0010839 | Not Fully Discharged, when range change from 200uA to 2mA
1 | 8 | 8 | 58 | 2 | F6 | C6 | 4 | 7 | 6B | 69 |  | 600 | 0.002 | 16172548 | -604668 | 0.00031716
1 | 8 | 8 | BC | 2 | F1 | E5 | 46 | 7 | BC | 93 |  | 700 | 0.002 | 15852870 | -924346 | 0.00048484
1 | 8 | 8 | 20 | 3 | E2 | D9 | 5E | 7 | 52 | 47 |  | 800 | 0.002 | 14866782 | -1910434 | 0.0010021
1 | 8 | 8 | 84 | 3 | D1 | 4 | 97 | 7 | 82 | 4D |  | 900 | 0.002 | 13698199 | -3079017 | 0.001615
1 | 8 | 8 | E8 | 3 | C8 | BD | 9A | 7 | 59 | 8 |  | 1000 | 0.002 | 13155738 | -3621478 | 0.0018995
1 | 8 | 8 | 4C | 4 | BE | D1 | 80 | 7 | 24 | D9 |  | 1100 | 0.002 | 12505472 | -4271744 | 0.0022406
1 | 8 | 8 | 4C | 4 | BE | D3 | E7 | 8 | EF | 2D |  | 1100 | 0.02 | 12506087 | -4271129 | 0.022403 | OverShoot even at Highest Range,Not yet Discharging
1 | 8 | 8 | B0 | 4 | F6 | 7 | 81 | 8 | 87 | 29 |  | 1200 | 0.02 | 16123777 | -653439 | 0.0034274
1 | 8 | 8 | 14 | 5 | F3 | E9 | 8C | 8 | C7 | A4 |  | 1300 | 0.02 | 15985036 | -792180 | 0.0041551
1 | 8 | 8 | 78 | 5 | F1 | BE | 3F | 8 | 0B | 90 |  | 1400 | 0.02 | 15842879 | -934337 | 0.0049008
1 | 8 | 8 | DC | 5 | EF | 96 | FE | 8 | C5 | 4 |  | 1500 | 0.02 | 15701758 | -1075458 | 0.005641
1 | 8 | 8 | 40 | 6 | ED | AB | E1 | 8 | 4 | DC |  | 1600 | 0.02 | 15576033 | -1201183 | 0.0063005
1 | 8 | 8 | A4 | 6 | EC | E9 | 4F | 8 | CE | B0 |  | 1700 | 0.02 | 15526223 | -1250993 | 0.0065617
1 | 8 | 8 | 8 | 7 | EA | B1 | C1 | 8 | 0F | E7 |  | 1800 | 0.02 | 15380929 | -1396287 | 0.0073238
1 | 8 | 8 | 6C | 7 | E8 | 64 | 39 | 8 | 54 | 43 |  | 1900 | 0.02 | 15230009 | -1547207 | 0.0081154
1 | 8 | 8 | D0 | 7 | E6 | 19 | 29 | 8 | D1 | 8F |  | 2000 | 0.02 | 15079721 | -1697495 | 0.0089037
1 | 8 | 8 | 34 | 8 | E4 | 0A | C9 | 8 | 2B | D7 |  | 2100 | 0.02 | 14944969 | -1832247 | 0.0096105
1 | 8 | 8 | 98 | 8 | E3 | 7D | 1B | 8 | DE | 75 |  | 2200 | 0.02 | 14908699 | -1868517 | 0.0098008
1 | 8 | 8 | FC | 8 | E1 | 30 | 5D | 8 | 75 | 9E |  | 2300 | 0.02 | 14757981 | -2019235 | 0.010591
1 | 8 | 8 | 60 | 9 | DE | E3 | 58 | 8 | AB | 7F |  | 2400 | 0.02 | 14607192 | -2170024 | 0.011382
1 | 8 | 8 | C4 | 9 | DC | 8F | 2A | 8 | 57 | 5E |  | 2500 | 0.02 | 14454570 | -2322646 | 0.012183
1 | 8 | 8 | 28 | 0A | DA | 47 | 7F | 8 | BB | D4 |  | 2600 | 0.02 | 14305151 | -2472065 | 0.012966
1 | 8 | 8 | 8C | 0A | D8 | 50 | 31 | 8 | 26 | EC |  | 2700 | 0.02 | 14176305 | -2600911 | 0.013642
1 | 8 | 8 | F0 | 0A | D7 | 98 | 5D | 8 | 82 | FA |  | 2800 | 0.02 | 14129245 | -2647971 | 0.013889
1 | 8 | 8 | 54 | 0B | D5 | 48 | 85 | 8 | FC | 5F |  | 2900 | 0.02 | 13977733 | -2799483 | 0.014684
1 | 8 | 8 | B8 | 0B | D2 | FF | 3 | 8 | 39 | C1 |  | 3000 | 0.02 | 13827843 | -2949373 | 0.01547
Test Results Using Diode and 150 ohm Series Resistor, Forward Bias (Dummy Specimen) Date : 11 May 2022
Hexadecimal Data From Device (Fetch Measurement) | [11 empty] | Bias Voltage (mV) | Measuement Range (A) | Unsigned Raw AD Value | Signed Raw AD Value | Resultant Bias Current (A)
1 | 8 | 8 | 0 | 0 | 0 | 4 | BD | 1 | 7E | 57 |  | 0 | 2e-09 | 1213 | 1213 | -6.3624e-13
1 | 8 | 8 | 64 | 0 | E4 | 1 | 9E | 1 | 49 | B2 |  | 100 | 2e-09 | 14942622 | -1834594 | 9.6228e-10
1 | 8 | 8 | C8 | 0 | F4 | DB | 38 | 5 | 0E | 86 |  | 200 | 2e-05 | 16046904 | -730312 | 3.8306e-06
1 | 8 | 8 | 2C | 1 | F9 | 3B | D8 | 5 | 6E | 38 |  | 300 | 2e-05 | 16333784 | -443432 | 2.3259e-06
1 | 8 | 8 | 90 | 1 | E8 | AD | E5 | 5 | 81 | 44 |  | 400 | 2e-05 | 15248869 | -1528347 | 8.0165e-06
1 | 8 | 8 | F4 | 1 | E0 | 78 | 25 | 7 | 4B | 39 |  | 500 | 0.002 | 14710821 | -2066395 | 0.0010839
1 | 8 | 8 | 58 | 2 | F6 | C6 | 4 | 7 | 6B | 69 |  | 600 | 0.002 | 16172548 | -604668 | 0.00031716
1 | 8 | 8 | BC | 2 | F1 | E5 | 46 | 7 | BC | 93 |  | 700 | 0.002 | 15852870 | -924346 | 0.00048484
1 | 8 | 8 | 20 | 3 | E2 | D9 | 5E | 7 | 52 | 47 |  | 800 | 0.002 | 14866782 | -1910434 | 0.0010021
1 | 8 | 8 | 84 | 3 | D1 | 4 | 97 | 7 | 82 | 4D |  | 900 | 0.002 | 13698199 | -3079017 | 0.001615
1 | 8 | 8 | E8 | 3 | C8 | BD | 9A | 7 | 59 | 8 |  | 1000 | 0.002 | 13155738 | -3621478 | 0.0018995
1 | 8 | 8 | 4C | 4 | BE | D3 | E7 | 8 | EF | 2D |  | 1100 | 0.02 | 12506087 | -4271129 | 0.022403
1 | 8 | 8 | B0 | 4 | F6 | 7 | 81 | 8 | 87 | 29 |  | 1200 | 0.02 | 16123777 | -653439 | 0.0034274
1 | 8 | 8 | 14 | 5 | F3 | E9 | 8C | 8 | C7 | A4 |  | 1300 | 0.02 | 15985036 | -792180 | 0.0041551
1 | 8 | 8 | 78 | 5 | F1 | BE | 3F | 8 | 0B | 90 |  | 1400 | 0.02 | 15842879 | -934337 | 0.0049008
1 | 8 | 8 | DC | 5 | EF | 96 | FE | 8 | C5 | 4 |  | 1500 | 0.02 | 15701758 | -1075458 | 0.005641
1 | 8 | 8 | 40 | 6 | ED | AB | E1 | 8 | 4 | DC |  | 1600 | 0.02 | 15576033 | -1201183 | 0.0063005
1 | 8 | 8 | A4 | 6 | EC | E9 | 4F | 8 | CE | B0 |  | 1700 | 0.02 | 15526223 | -1250993 | 0.0065617
1 | 8 | 8 | 8 | 7 | EA | B1 | C1 | 8 | 0F | E7 |  | 1800 | 0.02 | 15380929 | -1396287 | 0.0073238
1 | 8 | 8 | 6C | 7 | E8 | 64 | 39 | 8 | 54 | 43 |  | 1900 | 0.02 | 15230009 | -1547207 | 0.0081154
1 | 8 | 8 | D0 | 7 | E6 | 19 | 29 | 8 | D1 | 8F |  | 2000 | 0.02 | 15079721 | -1697495 | 0.0089037
1 | 8 | 8 | 34 | 8 | E4 | 0A | C9 | 8 | 2B | D7 |  | 2100 | 0.02 | 14944969 | -1832247 | 0.0096105
1 | 8 | 8 | 98 | 8 | E3 | 7D | 1B | 8 | DE | 75 |  | 2200 | 0.02 | 14908699 | -1868517 | 0.0098008
1 | 8 | 8 | FC | 8 | E1 | 30 | 5D | 8 | 75 | 9E |  | 2300 | 0.02 | 14757981 | -2019235 | 0.010591
1 | 8 | 8 | 60 | 9 | DE | E3 | 58 | 8 | AB | 7F |  | 2400 | 0.02 | 14607192 | -2170024 | 0.011382
1 | 8 | 8 | C4 | 9 | DC | 8F | 2A | 8 | 57 | 5E |  | 2500 | 0.02 | 14454570 | -2322646 | 0.012183
1 | 8 | 8 | 28 | 0A | DA | 47 | 7F | 8 | BB | D4 |  | 2600 | 0.02 | 14305151 | -2472065 | 0.012966
1 | 8 | 8 | 8C | 0A | D8 | 50 | 31 | 8 | 26 | EC |  | 2700 | 0.02 | 14176305 | -2600911 | 0.013642
1 | 8 | 8 | F0 | 0A | D7 | 98 | 5D | 8 | 82 | FA |  | 2800 | 0.02 | 14129245 | -2647971 | 0.013889
1 | 8 | 8 | 54 | 0B | D5 | 48 | 85 | 8 | FC | 5F |  | 2900 | 0.02 | 13977733 | -2799483 | 0.014684
1 | 8 | 8 | B8 | 0B | D2 | FF | 3 | 8 | 39 | C1 |  | 3000 | 0.02 | 13827843 | -2949373 | 0.01547
Problem 1 : there is time required to discharge to actual current whenever there is current range change. The larger the Gap between actual and Range might need long time to discharge, this issue will also cause inaccurate for the Autorange to pick correct Range to apply.

## 2RangeSearch-Flow3
[8 empty] | AD Full Scale Voltage (V) : | 3.3
[8 empty] | AD Reference Voltage Offset (V) : | 1.5
[8 empty] | AD Resolution (bits) : | 24
[8 empty] | Computational Constant : | 3813003.6364 | [6 empty] | Note:
Autorange in 2 Ranges up from low to high Range using Flow3
[8 empty] | Measurement Range | Multiplier | [6 empty] | If the reading reaches more than present range, will go up 2 ranges, or to the highest range possible if it can't go up 2 ranges.
[8 empty] | 1 | 2e-09 |  | (2nA) |  |  |  |  | The connector will take another reading, then decide whether it needs to continue going up in range or go down pick the right range based on the reading.
[8 empty] | 2 | 2e-08 |  | (20nA) |  |  |  |  | sweepdelay=100ms
[8 empty] | 3 | 2e-07 |  | (200nA)
[8 empty] | 4 | 2e-06 |  | (2uA)
[8 empty] | 5 | 2e-05 |  | (20uA)
[8 empty] | 6 | 0.0002 |  | (200uA)
[8 empty] | 7 | 0.002 |  | (2mA)
[8 empty] | 8 | 0.02 |  | (20mA)
Test Results Using Diode and 150 ohm Series Resistor, Forward Bias (Dummy Specimen)
Hexadecimal Data From Device (Fetch Measurement) | [11 empty] | Bias Voltage (mV) | Measuement Range (A) | Unsigned Raw AD Value | Signed Raw AD Value | Resultant Bias Current (A)
1 | 8 | 8 | 0 | 0 | FF | 49 | B2 | 1 | DB | A4 |  | 0 | 2e-09 | 16730546 | -46670 | 2.4479e-11
1 | 8 | 8 | 64 | 0 | BF | EF | AB | 1 | 2D | F3 |  | 100 | 2e-09 | 12578731 | -4198485 | 2.2022e-09
1 | 8 | 8 | 64 | 0 | D5 | 11 | 5D | 3 | 97 | BA |  | 100 | 2e-07 | 13963613 | -2813603 | 1.4758e-07
1 | 8 | 8 | C8 | 0 | C1 | 70 | 51 | 3 | DE | F8 |  | 200 | 2e-07 | 12677201 | -4100015 | 2.1505e-07
1 | 8 | 8 | C8 | 0 | DC | 2E | 93 | 5 | 68 | 24 |  | 200 | 2e-05 | 14429843 | -2347373 | 1.2312e-05
1 | 8 | 8 | 2C | 1 | F9 | 17 | BF | 5 | 85 | C1 |  | 300 | 2e-05 | 16324543 | -452673 | 2.3744e-06
1 | 8 | 8 | 90 | 1 | E9 | DE | 90 | 5 | 57 | F3 |  | 400 | 2e-05 | 15326864 | -1450352 | 7.6074e-06
1 | 8 | 8 | F4 | 1 | C4 | D1 | 92 | 5 | 67 | D8 |  | 500 | 2e-05 | 12898706 | -3878510 | 2.0344e-05
1 | 8 | 8 | F4 | 1 | BE | DD | D8 | 7 | 8 | 62 |  | 500 | 0.002 | 12508632 | -4268584 | 0.002239 | Overshoot, range change from 20uA to 2mA
1 | 8 | 8 | F4 | 1 | E6 | 79 | 59 | 8 | 7A | B5 |  | 500 | 0.02 | 15104345 | -1672871 | 0.0087746 | No Fully Discharge, when range change from 2mA to 20mA
1 | 8 | 8 | 58 | 2 | FF | 4C | 8C | 8 | 6F | 1B |  | 600 | 0.02 | 16731276 | -45940 | 0.00024096
1 | 8 | 8 | 58 | 2 | FE | E9 | 2F | 7 | 46 | 30 |  | 600 | 0.002 | 16705839 | -71377 | 3.7439e-05
1 | 8 | 8 | 58 | 2 | F6 | E5 | 61 | 6 | 70 | 33 |  | 600 | 0.0002 | 16180577 | -596639 | 3.1295e-05
1 | 8 | 8 | BC | 2 | C0 | CF | 7B | 6 | 42 | F7 |  | 700 | 0.0002 | 12636027 | -4141189 | 0.00021721
1 | 8 | 8 | BC | 2 | C0 | D1 | 56 | 8 | BE | 65 |  | 700 | 0.02 | 12636502 | -4140714 | 0.021719 | Overshoot even at Higher Range,Not yet Discharging
1 | 8 | 8 | 20 | 3 | FC | D5 | A3 | 8 | 94 | F8 |  | 800 | 0.02 | 16569763 | -207453 | 0.0010881
1 | 8 | 8 | 20 | 3 | FC | 1E | B6 | 7 | AB | 92 |  | 800 | 0.002 | 16522934 | -254282 | 0.00013338
1 | 8 | 8 | 20 | 3 | E0 | 6 | 75 | 6 | BD | 35 |  | 800 | 0.0002 | 14681717 | -2095499 | 0.00010991
1 | 8 | 8 | 84 | 3 | C0 | CF | 0A | 6 | 5F | DF |  | 900 | 0.0002 | 12635914 | -4141302 | 0.00021722
1 | 8 | 8 | 84 | 3 | C0 | D1 | 2D | 8 | A5 | ED |  | 900 | 0.02 | 12636461 | -4140755 | 0.021719 | Overshoot even at Higher Range,Not yet Discharging
1 | 8 | 8 | E8 | 3 | F9 | 86 | F8 | 8 | 4F | 5D |  | 1000 | 0.02 | 16353016 | -424200 | 0.002225
1 | 8 | 8 | 4C | 4 | F7 | 74 | B7 | 8 | 74 | 92 |  | 1100 | 0.02 | 16217271 | -559945 | 0.002937
1 | 8 | 8 | B0 | 4 | F5 | 4D | 68 | 8 | E9 | 2B |  | 1200 | 0.02 | 16076136 | -701080 | 0.0036773
1 | 8 | 8 | 14 | 5 | F3 | 55 | 5B | 8 | 59 | B0 |  | 1300 | 0.02 | 15947099 | -830117 | 0.0043541
1 | 8 | 8 | 78 | 5 | F2 | D8 | 81 | 8 | 9A | 6B |  | 1400 | 0.02 | 15915137 | -862079 | 0.0045218
1 | 8 | 8 | DC | 5 | F0 | 9F | BF | 8 | 22 | 82 |  | 1500 | 0.02 | 15769535 | -1007681 | 0.0052855
1 | 8 | 8 | 40 | 6 | EE | 61 | 44 | 8 | 5F | F6 |  | 1600 | 0.02 | 15622468 | -1154748 | 0.0060569
1 | 8 | 8 | A4 | 6 | EC | 17 | 50 | 8 | A7 | 70 |  | 1700 | 0.02 | 15472464 | -1304752 | 0.0068437
1 | 8 | 8 | 8 | 7 | E9 | DA | EA | 8 | 60 | 8F |  | 1800 | 0.02 | 15325930 | -1451286 | 0.0076123
1 | 8 | 8 | 6C | 7 | E7 | DF | 99 | 8 | 5F | 72 |  | 1900 | 0.02 | 15196057 | -1581159 | 0.0082935
1 | 8 | 8 | D0 | 7 | E7 | 33 | 7D | 8 | CF | 7B |  | 2000 | 0.02 | 15151997 | -1625219 | 0.0085246
1 | 8 | 8 | 34 | 8 | E4 | F4 | 61 | 8 | 35 | E7 |  | 2100 | 0.02 | 15004769 | -1772447 | 0.0092969
1 | 8 | 8 | 98 | 8 | E2 | AD | A8 | 8 | AB | 40 |  | 2200 | 0.02 | 14855592 | -1921624 | 0.010079
1 | 8 | 8 | FC | 8 | E0 | 65 | 11 | 8 | 50 | B2 |  | 2300 | 0.02 | 14705937 | -2071279 | 0.010864
1 | 8 | 8 | 60 | 9 | DE | 2F | 61 | 8 | 79 | 10 |  | 2400 | 0.02 | 14561121 | -2216095 | 0.011624
1 | 8 | 8 | C4 | 9 | DC | 31 | F5 | 8 | 6F | 4A |  | 2500 | 0.02 | 14430709 | -2346507 | 0.012308
1 | 8 | 8 | 28 | 0A | DB | 68 | 0E | 8 | AF | B1 |  | 2600 | 0.02 | 14379022 | -2398194 | 0.012579
1 | 8 | 8 | 8C | 0A | D9 | 19 | 2E | 8 | FE | F6 |  | 2700 | 0.02 | 14227758 | -2549458 | 0.013372
1 | 8 | 8 | F0 | 0A | D6 | C6 | 98 | 8 | B1 | 84 |  | 2800 | 0.02 | 14075544 | -2701672 | 0.014171
1 | 8 | 8 | 54 | 0B | D4 | 71 | BA | 8 | 3C | 5E |  | 2900 | 0.02 | 13922746 | -2854470 | 0.014972
1 | 8 | 8 | B8 | 0B | D2 | 72 | BE | 8 | D8 | BA |  | 3000 | 0.02 | 13791934 | -2985282 | 0.015658
Test Results Using Diode and 150 ohm Series Resistor, Forward Bias (Dummy Specimen)
Hexadecimal Data From Device (Fetch Measurement) | [11 empty] | Bias Voltage (mV) | Measuement Range (A) | Unsigned Raw AD Value | Signed Raw AD Value | Resultant Bias Current (A)
1 | 8 | 8 | 0 | 0 | FF | 49 | B2 | 1 | DB | A4 |  | 0 | 2e-09 | 16730546 | -46670 | 2.4479e-11
1 | 8 | 8 | 64 | 0 | D5 | 11 | 5D | 3 | 97 | BA |  | 100 | 2e-07 | 13963613 | -2813603 | 1.4758e-07
1 | 8 | 8 | C8 | 0 | DC | 2E | 93 | 5 | 68 | 24 |  | 200 | 2e-05 | 14429843 | -2347373 | 1.2312e-05
1 | 8 | 8 | 2C | 1 | F9 | 17 | BF | 5 | 85 | C1 |  | 300 | 2e-05 | 16324543 | -452673 | 2.3744e-06
1 | 8 | 8 | 90 | 1 | E9 | DE | 90 | 5 | 57 | F3 |  | 400 | 2e-05 | 15326864 | -1450352 | 7.6074e-06
1 | 8 | 8 | F4 | 1 | E6 | 79 | 59 | 8 | 7A | B5 |  | 500 | 0.02 | 15104345 | -1672871 | 0.0087746
1 | 8 | 8 | 58 | 2 | F6 | E5 | 61 | 6 | 70 | 33 |  | 600 | 0.0002 | 16180577 | -596639 | 3.1295e-05
1 | 8 | 8 | BC | 2 | C0 | D1 | 56 | 8 | BE | 65 |  | 700 | 0.02 | 12636502 | -4140714 | 0.021719
1 | 8 | 8 | 20 | 3 | E0 | 6 | 75 | 6 | BD | 35 |  | 800 | 0.0002 | 14681717 | -2095499 | 0.00010991
1 | 8 | 8 | 84 | 3 | C0 | D1 | 2D | 8 | A5 | ED |  | 900 | 0.02 | 12636461 | -4140755 | 0.021719
1 | 8 | 8 | E8 | 3 | F9 | 86 | F8 | 8 | 4F | 5D |  | 1000 | 0.02 | 16353016 | -424200 | 0.002225
1 | 8 | 8 | 4C | 4 | F7 | 74 | B7 | 8 | 74 | 92 |  | 1100 | 0.02 | 16217271 | -559945 | 0.002937
1 | 8 | 8 | B0 | 4 | F5 | 4D | 68 | 8 | E9 | 2B |  | 1200 | 0.02 | 16076136 | -701080 | 0.0036773
1 | 8 | 8 | 14 | 5 | F3 | 55 | 5B | 8 | 59 | B0 |  | 1300 | 0.02 | 15947099 | -830117 | 0.0043541
1 | 8 | 8 | 78 | 5 | F2 | D8 | 81 | 8 | 9A | 6B |  | 1400 | 0.02 | 15915137 | -862079 | 0.0045218
1 | 8 | 8 | DC | 5 | F0 | 9F | BF | 8 | 22 | 82 |  | 1500 | 0.02 | 15769535 | -1007681 | 0.0052855
1 | 8 | 8 | 40 | 6 | EE | 61 | 44 | 8 | 5F | F6 |  | 1600 | 0.02 | 15622468 | -1154748 | 0.0060569
1 | 8 | 8 | A4 | 6 | EC | 17 | 50 | 8 | A7 | 70 |  | 1700 | 0.02 | 15472464 | -1304752 | 0.0068437
1 | 8 | 8 | 8 | 7 | E9 | DA | EA | 8 | 60 | 8F |  | 1800 | 0.02 | 15325930 | -1451286 | 0.0076123
1 | 8 | 8 | 6C | 7 | E7 | DF | 99 | 8 | 5F | 72 |  | 1900 | 0.02 | 15196057 | -1581159 | 0.0082935
1 | 8 | 8 | D0 | 7 | E7 | 33 | 7D | 8 | CF | 7B |  | 2000 | 0.02 | 15151997 | -1625219 | 0.0085246
1 | 8 | 8 | 34 | 8 | E4 | F4 | 61 | 8 | 35 | E7 |  | 2100 | 0.02 | 15004769 | -1772447 | 0.0092969
1 | 8 | 8 | 98 | 8 | E2 | AD | A8 | 8 | AB | 40 |  | 2200 | 0.02 | 14855592 | -1921624 | 0.010079
1 | 8 | 8 | FC | 8 | E0 | 65 | 11 | 8 | 50 | B2 |  | 2300 | 0.02 | 14705937 | -2071279 | 0.010864
1 | 8 | 8 | 60 | 9 | DE | 2F | 61 | 8 | 79 | 10 |  | 2400 | 0.02 | 14561121 | -2216095 | 0.011624
1 | 8 | 8 | C4 | 9 | DC | 31 | F5 | 8 | 6F | 4A |  | 2500 | 0.02 | 14430709 | -2346507 | 0.012308
1 | 8 | 8 | 28 | 0A | DB | 68 | 0E | 8 | AF | B1 |  | 2600 | 0.02 | 14379022 | -2398194 | 0.012579
1 | 8 | 8 | 8C | 0A | D9 | 19 | 2E | 8 | FE | F6 |  | 2700 | 0.02 | 14227758 | -2549458 | 0.013372
1 | 8 | 8 | F0 | 0A | D6 | C6 | 98 | 8 | B1 | 84 |  | 2800 | 0.02 | 14075544 | -2701672 | 0.014171
1 | 8 | 8 | 54 | 0B | D4 | 71 | BA | 8 | 3C | 5E |  | 2900 | 0.02 | 13922746 | -2854470 | 0.014972
1 | 8 | 8 | B8 | 0B | D2 | 72 | BE | 8 | D8 | BA |  | 3000 | 0.02 | 13791934 | -2985282 | 0.015658
Problem 1: There is time required to discharge to actual current whenever there is current range change. The larger the Gap between actual and Range might need long time to discharge, this issue will also cause inaccurate for the Autorange to pick correct Range to apply.
Problem 2: Overshoot when range change,
No more range switch as this is the largest Range, and 20mA will be use as final range.

## 3RangeSearch-Flow3
[8 empty] | AD Full Scale Voltage (V) : | 3.3
[8 empty] | AD Reference Voltage Offset (V) : | 1.5
[8 empty] | AD Resolution (bits) : | 24
[8 empty] | Computational Constant : | 3813003.6364
[8 empty] | Measurement Range | Multiplier
[8 empty] | 1 | 2e-09 |  | (2nA)
[8 empty] | 2 | 2e-08 |  | (20nA) |  |  |  |  | Note:
[8 empty] | 3 | 2e-07 |  | (200nA) |  |  |  |  | Autorange in 3 Ranges up from low to high Range using Flow3
[8 empty] | 4 | 2e-06 |  | (2uA) |  |  |  |  | If the reading reaches more than present range, will go up 3 ranges, or to the highest range possible if it can't go up 3 ranges.
[8 empty] | 5 | 2e-05 |  | (20uA) |  |  |  |  | The connector will take another reading, then decide whether it needs to continue going up in range or go down pick the right range based on the reading.
[8 empty] | 6 | 0.0002 |  | (200uA) |  |  |  |  | sweepdelay=100ms
[8 empty] | 7 | 0.002 |  | (2mA)
[8 empty] | 8 | 0.02 |  | (20mA)
Test Results Using Diode and 150 ohm Series Resistor, Forward Bias (Dummy Specimen)
Hexadecimal Data From Device (Fetch Measurement) | [11 empty] | Bias Voltage (mV) | Measuement Range (A) | Unsigned Raw AD Value | Signed Raw AD Value | Resultant Bias Current (A)
1 | 8 | 8 | 0 | 0 | 0 | 4 | F2 | 1 | 4A | 67 |  | 0 | 2e-09 | 1266 | 1266 | -6.6404e-13
1 | 8 | 8 | 64 | 0 | D4 | 6C | D1 | 1 | E3 | 5F |  | 100 | 2e-09 | 13921489 | -2855727 | 1.4979e-09
1 | 8 | 8 | C8 | 0 | BE | DD | D4 | 1 | B5 | 9C |  | 200 | 2e-09 | 12508628 | -4268588 | 2.239e-09
1 | 8 | 8 | C8 | 0 | BE | DD | A4 | 4 | 50 | 5F |  | 200 | 2e-06 | 12508580 | -4268636 | 2.239e-06
1 | 8 | 8 | C8 | 0 | D8 | 9B | 91 | 7 | F8 | 53 |  | 200 | 0.002 | 14195601 | -2581615 | 0.0013541
1 | 8 | 8 | 2C | 1 | FD | BD | 6A | 7 | 7B | 80 |  | 300 | 0.002 | 16629098 | -148118 | 7.7691e-05
1 | 8 | 8 | 2C | 1 | FF | 33 | C8 | 6 | A2 | B3 |  | 300 | 0.0002 | 16724936 | -52280 | 2.7422e-06
1 | 8 | 8 | 2C | 1 | FE | ED | 8 | 5 | D3 | 74 |  | 300 | 2e-05 | 16706824 | -70392 | 3.6922e-07
1 | 8 | 8 | 2C | 1 | F9 | 1C | 5D | 4 | 7D | 63 |  | 300 | 2e-06 | 16325725 | -451491 | 2.3682e-07
1 | 8 | 8 | 90 | 1 | BF | 30 | 6F | 4 | A6 | 7E |  | 400 | 2e-06 | 12529775 | -4247441 | 2.2279e-06
1 | 8 | 8 | 90 | 1 | BE | DD | DC | 7 | 2 | 86 |  | 400 | 0.002 | 12508636 | -4268580 | 0.002239 | Overshoot when range change from 2uA to 2mA,Not yet Discharging
1 | 8 | 8 | 90 | 1 | DB | 39 | 75 | 8 | 62 | 29 |  | 400 | 0.02 | 14367093 | -2410123 | 0.012642 | No Fully Discharge when Range change from 2mA to 20mA
1 | 8 | 8 | F4 | 1 | FF | CE | 4D | 8 | C2 | CF |  | 500 | 0.02 | 16764493 | -12723 | 6.6735e-05
1 | 8 | 8 | F4 | 1 | FF | B3 | 1E | 6 | EF | E3 |  | 500 | 0.0002 | 16757534 | -19682 | 1.0324e-06
1 | 8 | 8 | F4 | 1 | E0 | 15 | CB | 4 | D7 | 45 |  | 500 | 2e-06 | 14685643 | -2091573 | 1.0971e-06
1 | 8 | 8 | 58 | 2 | BE | DD | D5 | 4 | 10 | 9F |  | 600 | 2e-06 | 12508629 | -4268587 | 2.239e-06
1 | 8 | 8 | 58 | 2 | BE | DD | D7 | 7 | 51 | FE |  | 600 | 0.002 | 12508631 | -4268585 | 0.002239 | Overshoot when range change from 2uA to 2mA,Not yet Discharging
1 | 8 | 8 | 58 | 2 | CD | E3 | 5 | 8 | 36 | 12 |  | 600 | 0.02 | 13492997 | -3284219 | 0.017226 | No Fully Discharge when Range change from 2mA to 20mA
1 | 8 | 8 | BC | 2 | FE | 50 | 7A | 8 | FF | 65 |  | 700 | 0.02 | 16666746 | -110470 | 0.00057944
1 | 8 | 8 | BC | 2 | FD | A9 | 43 | 7 | 7D | 44 |  | 700 | 0.002 | 16623939 | -153277 | 8.0397e-05
1 | 8 | 8 | BC | 2 | EC | D6 | 5B | 6 | 82 | 60 |  | 700 | 0.0002 | 15521371 | -1255845 | 6.5872e-05
1 | 8 | 8 | 20 | 3 | C0 | CF | 9A | 6 | 2A | FB |  | 800 | 0.0002 | 12636058 | -4141158 | 0.00021721
1 | 8 | 8 | 20 | 3 | C0 | D1 | 63 | 8 | 89 | 69 |  | 800 | 0.02 | 12636515 | -4140701 | 0.021719
1 | 8 | 8 | 84 | 3 | FB | 5A | 7F | 8 | E5 | 83 |  | 900 | 0.02 | 16472703 | -304513 | 0.0015972
1 | 8 | 8 | 84 | 3 | FA | 7A | 32 | 7 | 90 | E1 |  | 900 | 0.002 | 16415282 | -361934 | 0.00018984
1 | 8 | 8 | 84 | 3 | D3 | 7A | 2A | 6 | 53 | 7D |  | 900 | 0.0002 | 13859370 | -2917846 | 0.00015305
1 | 8 | 8 | E8 | 3 | C0 | CF | 0F | 6 | 55 | E3 |  | 1000 | 0.0002 | 12635919 | -4141297 | 0.00021722
1 | 8 | 8 | E8 | 3 | C0 | D1 | 0D | 8 | B5 | 41 |  | 1000 | 0.02 | 12636429 | -4140787 | 0.021719
1 | 8 | 8 | 4C | 4 | F7 | E7 | A9 | 8 | 8D | 1F |  | 1100 | 0.02 | 16246697 | -530519 | 0.0027827
1 | 8 | 8 | B0 | 4 | F5 | BD | D8 | 8 | 9C | D8 |  | 1200 | 0.02 | 16104920 | -672296 | 0.0035263
1 | 8 | 8 | 14 | 5 | F3 | 9A | 32 | 8 | 47 | DF |  | 1300 | 0.02 | 15964722 | -812494 | 0.0042617
1 | 8 | 8 | 78 | 5 | F1 | 78 | B3 | 8 | 8F | 6D |  | 1400 | 0.02 | 15825075 | -952141 | 0.0049942
1 | 8 | 8 | DC | 5 | EF | 93 | 6D | 8 | B9 | F5 |  | 1500 | 0.02 | 15700845 | -1076371 | 0.0056458
1 | 8 | 8 | 40 | 6 | EE | DA | 12 | 8 | 10 | 73 |  | 1600 | 0.02 | 15653394 | -1123822 | 0.0058947
1 | 8 | 8 | A4 | 6 | EC | 98 | EC | 8 | E6 | 5B |  | 1700 | 0.02 | 15505644 | -1271572 | 0.0066697
1 | 8 | 8 | 8 | 7 | EA | 5B | 85 | 8 | 1D | 13 |  | 1800 | 0.02 | 15358853 | -1418363 | 0.0074396
1 | 8 | 8 | 6C | 7 | E8 | 0F | 74 | 8 | 10 | CF |  | 1900 | 0.02 | 15208308 | -1568908 | 0.0082292
1 | 8 | 8 | D0 | 7 | E5 | F3 | 80 | 8 | 8E | 6F |  | 2000 | 0.02 | 15070080 | -1707136 | 0.0089543
1 | 8 | 8 | 34 | 8 | E5 | 7D | C9 | 8 | 9A | 31 |  | 2100 | 0.02 | 15039945 | -1737271 | 0.0091123
1 | 8 | 8 | 98 | 8 | E3 | 3C | 99 | 8 | EE | C1 |  | 2200 | 0.02 | 14892185 | -1885031 | 0.0098874
1 | 8 | 8 | FC | 8 | E0 | F2 | A3 | 8 | 95 | FE |  | 2300 | 0.02 | 14742179 | -2035037 | 0.010674
1 | 8 | 8 | 60 | 9 | DE | AE | 30 | 8 | 14 | A8 |  | 2400 | 0.02 | 14593584 | -2183632 | 0.011454
1 | 8 | 8 | C4 | 9 | DC | 5C | 95 | 8 | D6 | 97 |  | 2500 | 0.02 | 14441621 | -2335595 | 0.012251
1 | 8 | 8 | 28 | 0A | DA | 43 | F6 | 8 | 9D | 85 |  | 2600 | 0.02 | 14304246 | -2472970 | 0.012971
1 | 8 | 8 | 8C | 0A | D9 | C9 | 27 | 8 | F9 | 5F |  | 2700 | 0.02 | 14272807 | -2504409 | 0.013136
1 | 8 | 8 | F0 | 0A | D7 | 75 | 3C | 8 | 3B | 5F |  | 2800 | 0.02 | 14120252 | -2656964 | 0.013936
1 | 8 | 8 | 54 | 0B | D5 | 27 | 70 | 8 | 8B | 12 |  | 2900 | 0.02 | 13969264 | -2807952 | 0.014728
1 | 8 | 8 | B8 | 0B | D2 | DA | 3E | 8 | 38 | 9A |  | 3000 | 0.02 | 13818430 | -2958786 | 0.015519
Test Results Using Diode and 150 ohm Series Resistor, Forward Bias (Dummy Specimen)
Hexadecimal Data From Device (Fetch Measurement) | [11 empty] | Bias Voltage (mV) | Measuement Range (A) | Unsigned Raw AD Value | Signed Raw AD Value | Resultant Bias Current (A)
1 | 8 | 8 | 0 | 0 | 0 | 4 | F2 | 1 | 4A | 67 |  | 0 | 2e-09 | 1266 | 1266 | -6.6404e-13
1 | 8 | 8 | 64 | 0 | D4 | 6C | D1 | 1 | E3 | 5F |  | 100 | 2e-09 | 13921489 | -2855727 | 1.4979e-09
1 | 8 | 8 | C8 | 0 | D8 | 9B | 91 | 7 | F8 | 53 |  | 200 | 0.002 | 14195601 | -2581615 | 0.0013541
1 | 8 | 8 | 2C | 1 | F9 | 1C | 5D | 4 | 7D | 63 |  | 300 | 2e-06 | 16325725 | -451491 | 2.3682e-07
1 | 8 | 8 | 90 | 1 | DB | 39 | 75 | 8 | 62 | 29 |  | 400 | 0.02 | 14367093 | -2410123 | 0.012642
1 | 8 | 8 | F4 | 1 | E0 | 15 | CB | 4 | D7 | 45 |  | 500 | 2e-06 | 14685643 | -2091573 | 1.0971e-06
1 | 8 | 8 | 58 | 2 | CD | E3 | 5 | 8 | 36 | 12 |  | 600 | 0.02 | 13492997 | -3284219 | 0.017226
1 | 8 | 8 | BC | 2 | EC | D6 | 5B | 6 | 82 | 60 |  | 700 | 0.0002 | 15521371 | -1255845 | 6.5872e-05
1 | 8 | 8 | 20 | 3 | C0 | D1 | 63 | 8 | 89 | 69 |  | 800 | 0.02 | 12636515 | -4140701 | 0.021719
1 | 8 | 8 | 84 | 3 | D3 | 7A | 2A | 6 | 53 | 7D |  | 900 | 0.0002 | 13859370 | -2917846 | 0.00015305
1 | 8 | 8 | E8 | 3 | C0 | D1 | 0D | 8 | B5 | 41 |  | 1000 | 0.02 | 12636429 | -4140787 | 0.021719
1 | 8 | 8 | 4C | 4 | F7 | E7 | A9 | 8 | 8D | 1F |  | 1100 | 0.02 | 16246697 | -530519 | 0.0027827
1 | 8 | 8 | B0 | 4 | F5 | BD | D8 | 8 | 9C | D8 |  | 1200 | 0.02 | 16104920 | -672296 | 0.0035263
1 | 8 | 8 | 14 | 5 | F3 | 9A | 32 | 8 | 47 | DF |  | 1300 | 0.02 | 15964722 | -812494 | 0.0042617
1 | 8 | 8 | 78 | 5 | F1 | 78 | B3 | 8 | 8F | 6D |  | 1400 | 0.02 | 15825075 | -952141 | 0.0049942
1 | 8 | 8 | DC | 5 | EF | 93 | 6D | 8 | B9 | F5 |  | 1500 | 0.02 | 15700845 | -1076371 | 0.0056458
1 | 8 | 8 | 40 | 6 | EE | DA | 12 | 8 | 10 | 73 |  | 1600 | 0.02 | 15653394 | -1123822 | 0.0058947
1 | 8 | 8 | A4 | 6 | EC | 98 | EC | 8 | E6 | 5B |  | 1700 | 0.02 | 15505644 | -1271572 | 0.0066697
1 | 8 | 8 | 8 | 7 | EA | 5B | 85 | 8 | 1D | 13 |  | 1800 | 0.02 | 15358853 | -1418363 | 0.0074396
1 | 8 | 8 | 6C | 7 | E8 | 0F | 74 | 8 | 10 | CF |  | 1900 | 0.02 | 15208308 | -1568908 | 0.0082292
1 | 8 | 8 | D0 | 7 | E5 | F3 | 80 | 8 | 8E | 6F |  | 2000 | 0.02 | 15070080 | -1707136 | 0.0089543
1 | 8 | 8 | 34 | 8 | E5 | 7D | C9 | 8 | 9A | 31 |  | 2100 | 0.02 | 15039945 | -1737271 | 0.0091123
1 | 8 | 8 | 98 | 8 | E3 | 3C | 99 | 8 | EE | C1 |  | 2200 | 0.02 | 14892185 | -1885031 | 0.0098874
1 | 8 | 8 | FC | 8 | E0 | F2 | A3 | 8 | 95 | FE |  | 2300 | 0.02 | 14742179 | -2035037 | 0.010674
1 | 8 | 8 | 60 | 9 | DE | AE | 30 | 8 | 14 | A8 |  | 2400 | 0.02 | 14593584 | -2183632 | 0.011454
1 | 8 | 8 | C4 | 9 | DC | 5C | 95 | 8 | D6 | 97 |  | 2500 | 0.02 | 14441621 | -2335595 | 0.012251
1 | 8 | 8 | 28 | 0A | DA | 43 | F6 | 8 | 9D | 85 |  | 2600 | 0.02 | 14304246 | -2472970 | 0.012971
1 | 8 | 8 | 8C | 0A | D9 | C9 | 27 | 8 | F9 | 5F |  | 2700 | 0.02 | 14272807 | -2504409 | 0.013136
1 | 8 | 8 | F0 | 0A | D7 | 75 | 3C | 8 | 3B | 5F |  | 2800 | 0.02 | 14120252 | -2656964 | 0.013936
1 | 8 | 8 | 54 | 0B | D5 | 27 | 70 | 8 | 8B | 12 |  | 2900 | 0.02 | 13969264 | -2807952 | 0.014728
1 | 8 | 8 | B8 | 0B | D2 | DA | 3E | 8 | 38 | 9A |  | 3000 | 0.02 | 13818430 | -2958786 | 0.015519
Problem 1: there is time required to discharge to actual current whenever there is current range change. The larger the Gap between actual and Range might need long time to discharge, this issue will also cause inaccurate for the Autorange to pick correct Range to apply.
Problem 2: Overshoot when range change, It required time to discharge
No more range switch as this is the largest Range and 20mA will be use as the final range.
As the current discharge along the time, the most sensitive range should be apply at 2mA instead of 20mA

## Fixed20mA
[8 empty] | AD Full Scale Voltage (V) : | 3.3
[8 empty] | AD Reference Voltage Offset (V) : | 1.5
[8 empty] | AD Resolution (bits) : | 24
[8 empty] | Computational Constant : | 3813003.6364
Note:
[8 empty] | Measurement Range | Multiplier | [6 empty] | Fixed Range in 20mA
[8 empty] | 1 | 2e-09 |  | (2nA) |  |  |  |  | sweepdelay=100ms
[8 empty] | 2 | 2e-08 |  | (20nA)
[8 empty] | 3 | 2e-07 |  | (200nA)
[8 empty] | 4 | 2e-06 |  | (2uA)
[8 empty] | 5 | 2e-05 |  | (20uA)
[8 empty] | 6 | 0.0002 |  | (200uA)
[8 empty] | 7 | 0.002 |  | (2mA)
[8 empty] | 8 | 0.02 |  | (20mA)
Test Results Using Diode and 150 ohm Series Resistor, Forward Bias (Dummy Specimen) Date : 11 May 2022
Hexadecimal Data From Device (Fetch Measurement) | [11 empty] | Bias Voltage (mV) | Measuement Range (A) | Unsigned Raw AD Value | Signed Raw AD Value | Resultant Bias Current (A)
1 | 8 | 8 | 0 | 0 | 0 | 4 | A9 | 8 | B1 | 51 |  | 0 | 0.02 | 1193 | 1193 | -6.2575e-06
1 | 8 | 8 | 64 | 0 | 0 | 4 | A9 | 8 | B9 | 75 |  | 100 | 0.02 | 1193 | 1193 | -6.2575e-06
1 | 8 | 8 | C8 | 0 | 0 | 4 | 73 | 8 | FA | 79 |  | 200 | 0.02 | 1139 | 1139 | -5.9743e-06
1 | 8 | 8 | 2C | 1 | 0 | 2 | F3 | 8 | 51 | 9C |  | 300 | 0.02 | 755 | 755 | -3.9601e-06
1 | 8 | 8 | 90 | 1 | FF | F7 | B4 | 8 | 58 | B6 |  | 400 | 0.02 | 16775092 | -2124 | 1.1141e-05
1 | 8 | 8 | F4 | 1 | FF | B4 | DD | 8 | 8F | 16 |  | 500 | 0.02 | 16757981 | -19235 | 0.00010089
1 | 8 | 8 | 58 | 2 | FE | EC | D6 | 8 | 54 | 65 |  | 600 | 0.02 | 16706774 | -70442 | 0.00036948
1 | 8 | 8 | BC | 2 | FD | AD | 9E | 8 | 25 | D1 |  | 700 | 0.02 | 16625054 | -152162 | 0.00079812
1 | 8 | 8 | 20 | 3 | FC | 24 | 1B | 8 | B7 | 0B |  | 800 | 0.02 | 16524315 | -252901 | 0.0013265
1 | 8 | 8 | 84 | 3 | FA | 96 | C4 | 8 | 56 | B0 |  | 900 | 0.02 | 16422596 | -354620 | 0.0018601
1 | 8 | 8 | E8 | 3 | F8 | C6 | 6 | 8 | 0F | 15 |  | 1000 | 0.02 | 16303622 | -473594 | 0.0024841
1 | 8 | 8 | 4C | 4 | F6 | FD | 1C | 8 | DB | B4 |  | 1100 | 0.02 | 16186652 | -590564 | 0.0030976
1 | 8 | 8 | B0 | 4 | F5 | 6F | AF | 8 | 1B | 11 |  | 1200 | 0.02 | 16084911 | -692305 | 0.0036313
1 | 8 | 8 | 14 | 5 | F3 | 5B | 53 | 8 | 3F | B3 |  | 1300 | 0.02 | 15948627 | -828589 | 0.0043461
1 | 8 | 8 | 78 | 5 | F1 | 7E | AA | 8 | 64 | FC |  | 1400 | 0.02 | 15826602 | -950614 | 0.0049862
1 | 8 | 8 | DC | 5 | EF | F1 | 37 | 8 | 22 | 8B |  | 1500 | 0.02 | 15724855 | -1052361 | 0.0055199
1 | 8 | 8 | 40 | 6 | ED | AE | 40 | 8 | 6D | 4D |  | 1600 | 0.02 | 15576640 | -1200576 | 0.0062973
1 | 8 | 8 | A4 | 6 | EB | C2 | 30 | 8 | 9F | FC |  | 1700 | 0.02 | 15450672 | -1326544 | 0.006958
1 | 8 | 8 | 8 | 7 | EA | 44 | 27 | 8 | 55 | B5 |  | 1800 | 0.02 | 15352871 | -1424345 | 0.007471
1 | 8 | 8 | 6C | 7 | E7 | E5 | A0 | 8 | 6D | 2F |  | 1900 | 0.02 | 15197600 | -1579616 | 0.0082854
1 | 8 | 8 | D0 | 7 | E5 | F9 | 10 | 8 | C2 | 6D |  | 2000 | 0.02 | 15071504 | -1705712 | 0.0089468
1 | 8 | 8 | 34 | 8 | E4 | 0D | 16 | 8 | C2 | 26 |  | 2100 | 0.02 | 14945558 | -1831658 | 0.0096074
1 | 8 | 8 | 98 | 8 | E2 | 1F | 50 | 8 | 48 | A7 |  | 2200 | 0.02 | 14819152 | -1958064 | 0.01027
1 | 8 | 8 | FC | 8 | E0 | 33 | 4A | 8 | 8B | 92 |  | 2300 | 0.02 | 14693194 | -2084022 | 0.010931
1 | 8 | 8 | 60 | 9 | DE | 34 | DB | 8 | 7A | 77 |  | 2400 | 0.02 | 14562523 | -2214693 | 0.011617
1 | 8 | 8 | C4 | 9 | DC | 37 | DB | 8 | 92 | EB |  | 2500 | 0.02 | 14432219 | -2344997 | 0.0123
1 | 8 | 8 | 28 | 0A | DA | 6B | A8 | 8 | 25 | ED |  | 2600 | 0.02 | 14314408 | -2462808 | 0.012918
1 | 8 | 8 | 8C | 0A | D8 | 51 | F5 | 8 | 25 | EC |  | 2700 | 0.02 | 14176757 | -2600459 | 0.01364
1 | 8 | 8 | F0 | 0A | D6 | 5B | A4 | 8 | 31 | 6A |  | 2800 | 0.02 | 14048164 | -2729052 | 0.014314
1 | 8 | 8 | 54 | 0B | D4 | A9 | 6 | 8 | CC | A5 |  | 2900 | 0.02 | 13936902 | -2840314 | 0.014898
1 | 8 | 8 | B8 | 0B | D2 | 77 | D1 | 8 | E5 | 4B |  | 3000 | 0.02 | 13793233 | -2983983 | 0.015652

## 20mAto2nA_RemPrev
[8 empty] | AD Full Scale Voltage (V) : | 3.3
[8 empty] | AD Reference Voltage Offset (V) : | 1.5
[8 empty] | AD Resolution (bits) : | 24
[8 empty] | Computational Constant : | 3813003.6364
[8 empty] | Measurement Range | Multiplier
[8 empty] | 1 | 2e-09 |  | (2nA)
[8 empty] | 2 | 2e-08 |  | (20nA) |  |  |  |  | Note:
[8 empty] | 3 | 2e-07 |  | (200nA) |  |  |  |  | Auto Range from High to low range (20mA-2nA)
[8 empty] | 4 | 2e-06 |  | (2uA) |  |  |  |  | Stay at previous Current Range for next voltage step
[8 empty] | 5 | 2e-05 |  | (20uA) |  |  |  |  | sweepdelay=100ms
[8 empty] | 6 | 0.0002 |  | (200uA)
[8 empty] | 7 | 0.002 |  | (2mA)
[8 empty] | 8 | 0.02 |  | (20mA)
Test Results Using Diode and 150 ohm Series Resistor, Forward Bias (Dummy Specimen) Date : 11 May 2022
Hexadecimal Data From Device (Fetch Measurement) | [11 empty] | Bias Voltage (mV) | Measuement Range (A) | Unsigned Raw AD Value | Signed Raw AD Value | Resultant Bias Current (A)
1 | 8 | 8 | 0 | 0 | FD | BA | 13 | 8 | 93 | B9 |  | 0 | 0.02 | 16628243 | -148973 | 0.00078139
1 | 8 | 8 | 0 | 0 | 0 | 4 | B8 | 7 | FD | 5 |  | 0 | 0.002 | 1208 | 1208 | -6.3362e-07
1 | 8 | 8 | 0 | 0 | FF | FF | 90 | 1 | 22 | E2 |  | 0 | 2e-09 | 16777104 | -112 | 5.8746e-14
1 | 8 | 8 | 64 | 0 | E2 | 5 | 67 | 1 | 4A | AB |  | 100 | 2e-09 | 14812519 | -1964697 | 1.0305e-09
1 | 8 | 8 | C8 | 0 | BE | DD | CE | 1 | BE | FC |  | 200 | 2e-09 | 12508622 | -4268594 | 2.239e-09
1 | 8 | 8 | C8 | 0 | BE | DD | CB | 2 | FD | AD |  | 200 | 2e-08 | 12508619 | -4268597 | 2.239e-08
1 | 8 | 8 | C8 | 0 | BE | DD | C7 | 3 | 39 | 6D |  | 200 | 2e-07 | 12508615 | -4268601 | 2.239e-07
1 | 8 | 8 | C8 | 0 | C0 | CE | F9 | 4 | 81 | 22 |  | 200 | 2e-06 | 12635897 | -4141319 | 2.1722e-06
1 | 8 | 8 | C8 | 0 | F8 | C8 | 92 | 5 | 82 | B3 |  | 200 | 2e-05 | 16304274 | -472942 | 2.4807e-06
1 | 8 | 8 | 2C | 1 | FC | E0 | AD | 5 | 38 | 5F |  | 300 | 2e-05 | 16572589 | -204627 | 1.0733e-06
1 | 8 | 8 | 2C | 1 | F9 | 59 | 81 | 4 | 34 | 76 |  | 300 | 2e-06 | 16341377 | -435839 | 2.2861e-07
1 | 8 | 8 | 90 | 1 | C1 | 12 | 92 | 4 | 5E | CC |  | 400 | 2e-06 | 12653202 | -4124014 | 2.1631e-06
1 | 8 | 8 | 90 | 1 | BE | DD | D0 | 5 | 86 | 47 |  | 400 | 2e-05 | 12508624 | -4268592 | 2.239e-05
1 | 8 | 8 | 90 | 1 | EB | F9 | E8 | 6 | 84 | 41 |  | 400 | 0.0002 | 15464936 | -1312280 | 6.8832e-05
1 | 8 | 8 | F4 | 1 | F2 | B8 | 88 | 6 | F3 | 2D |  | 500 | 0.0002 | 15906952 | -870264 | 4.5647e-05
1 | 8 | 8 | 58 | 2 | D1 | 6 | D4 | 6 | FD | E1 |  | 600 | 0.0002 | 13698772 | -3078444 | 0.00016147
1 | 8 | 8 | BC | 2 | C0 | D0 | 14 | 6 | 5E | C1 |  | 700 | 0.0002 | 12636180 | -4141036 | 0.00021721
1 | 8 | 8 | BC | 2 | C0 | EE | A4 | 7 | 8B | 0D |  | 700 | 0.002 | 12644004 | -4133212 | 0.002168 | OverShoot Range,Not yet Discharging
1 | 8 | 8 | BC | 2 | D7 | 56 | 16 | 8 | 3A | 38 |  | 700 | 0.02 | 14112278 | -2664938 | 0.013978 | Not Fully Discharged, when 2mA to 20mA
1 | 8 | 8 | 20 | 3 | FC | AC | 0B | 8 | 3A | E1 |  | 800 | 0.02 | 16559115 | -218101 | 0.001144
1 | 8 | 8 | 20 | 3 | FC | 24 | 15 | 7 | F3 | 6F |  | 800 | 0.002 | 16524309 | -252907 | 0.00013266
1 | 8 | 8 | 20 | 3 | DF | 3E | 0D | 6 | 12 | EC |  | 800 | 0.0002 | 14630413 | -2146803 | 0.0001126
1 | 8 | 8 | 84 | 3 | C0 | CF | EA | 6 | 16 | 1F |  | 900 | 0.0002 | 12636138 | -4141078 | 0.00021721
1 | 8 | 8 | 84 | 3 | C0 | D2 | 79 | 7 | 2B | 29 |  | 900 | 0.002 | 12636793 | -4140423 | 0.0021717 | OverShoot Range,Not yet Discharging
1 | 8 | 8 | 84 | 3 | BF | 31 | DB | 8 | FA | 6F |  | 900 | 0.02 | 12530139 | -4247077 | 0.022277 | Not Fully Discharged, when 2mA to 20mA
1 | 8 | 8 | E8 | 3 | F9 | 99 | 1A | 8 | 36 | 3B |  | 1000 | 0.02 | 16357658 | -419558 | 0.0022007
1 | 8 | 8 | 4C | 4 | F7 | 89 | A7 | 8 | E8 | A2 |  | 1100 | 0.02 | 16222631 | -554585 | 0.0029089
1 | 8 | 8 | B0 | 4 | F5 | 6D | 35 | 8 | D0 | 71 |  | 1200 | 0.02 | 16084277 | -692939 | 0.0036346
1 | 8 | 8 | 14 | 5 | F3 | 5B | 92 | 8 | 6E | 23 |  | 1300 | 0.02 | 15948690 | -828526 | 0.0043458
1 | 8 | 8 | 78 | 5 | F1 | 7E | AC | 8 | 67 | 5C |  | 1400 | 0.02 | 15826604 | -950612 | 0.0049862
1 | 8 | 8 | DC | 5 | F0 | BD | 57 | 8 | CC | 88 |  | 1500 | 0.02 | 15777111 | -1000105 | 0.0052458
1 | 8 | 8 | 40 | 6 | EE | 78 | 8A | 8 | DA | 51 |  | 1600 | 0.02 | 15628426 | -1148790 | 0.0060256
1 | 8 | 8 | A4 | 6 | EC | 38 | 69 | 8 | 84 | E9 |  | 1700 | 0.02 | 15480937 | -1296279 | 0.0067993
1 | 8 | 8 | 8 | 7 | E9 | FE | 2A | 8 | 70 | 84 |  | 1800 | 0.02 | 15334954 | -1442262 | 0.007565
1 | 8 | 8 | 6C | 7 | E7 | E5 | C4 | 8 | 47 | EF |  | 1900 | 0.02 | 15197636 | -1579580 | 0.0082852
1 | 8 | 8 | D0 | 7 | E7 | 58 | 20 | 8 | 86 | 37 |  | 2000 | 0.02 | 15161376 | -1615840 | 0.0084754
1 | 8 | 8 | 34 | 8 | E5 | 10 | A5 | 8 | 26 | EC |  | 2100 | 0.02 | 15012005 | -1765211 | 0.0092589
1 | 8 | 8 | 98 | 8 | E2 | CE | 2B | 8 | 3A | 6E |  | 2200 | 0.02 | 14863915 | -1913301 | 0.010036
1 | 8 | 8 | FC | 8 | E0 | 7C | 68 | 8 | A2 | E5 |  | 2300 | 0.02 | 14711912 | -2065304 | 0.010833
1 | 8 | 8 | 60 | 9 | DE | 35 | CC | 8 | 24 | 47 |  | 2400 | 0.02 | 14562764 | -2214452 | 0.011615
1 | 8 | 8 | C4 | 9 | DC | 38 | 8B | 8 | 9E | E8 |  | 2500 | 0.02 | 14432395 | -2344821 | 0.012299
1 | 8 | 8 | 28 | 0A | DB | 8A | DE | 8 | 52 | 47 |  | 2600 | 0.02 | 14387934 | -2389282 | 0.012532
1 | 8 | 8 | 8C | 0A | D9 | 3E | 11 | 8 | 5F | 0D |  | 2700 | 0.02 | 14237201 | -2540015 | 0.013323
1 | 8 | 8 | F0 | 0A | D6 | EC | 69 | 8 | D5 | DC |  | 2800 | 0.02 | 14085225 | -2691991 | 0.01412
1 | 8 | 8 | 54 | 0B | D4 | 9E | 18 | 8 | 74 | CB |  | 2900 | 0.02 | 13934104 | -2843112 | 0.014913
1 | 8 | 8 | B8 | 0B | D2 | 79 | 68 | 8 | F7 | 18 |  | 3000 | 0.02 | 13793640 | -2983576 | 0.015649

## 20mAto2nA
[8 empty] | AD Full Scale Voltage (V) : | 3.3
[8 empty] | AD Reference Voltage Offset (V) : | 1.5
[8 empty] | AD Resolution (bits) : | 24
[8 empty] | Computational Constant : | 3813003.6364
[8 empty] | Measurement Range | Multiplier
[8 empty] | 1 | 2e-09 |  | (2nA) |  |  |  |  | Note:
[8 empty] | 2 | 2e-08 |  | (20nA) |  |  |  |  | Auto Range from High to low range (20mA-2nA)
[8 empty] | 3 | 2e-07 |  | (200nA) |  |  |  |  | Every Voltage Step Start from (20mA)
[8 empty] | 4 | 2e-06 |  | (2uA) |  |  |  |  | sweepdelay=100ms
[8 empty] | 5 | 2e-05 |  | (20uA)
[8 empty] | 6 | 0.0002 |  | (200uA)
[8 empty] | 7 | 0.002 |  | (2mA)
[8 empty] | 8 | 0.02 |  | (20mA)
Test Results Using Diode and 150 ohm Series Resistor, Forward Bias (Dummy Specimen) Date : 11 May 2022
Hexadecimal Data From Device (Fetch Measurement) | [11 empty] | Bias Voltage (mV) | Measuement Range (A) | Unsigned Raw AD Value | Signed Raw AD Value | Resultant Bias Current (A)
1 | 8 | 8 | 0 | 0 | 0 | 4 | AE | 8 | B3 | 61 |  | 0 | 0.02 | 1198 | 1198 | -6.2838e-06
1 | 8 | 8 | 0 | 0 | 0 | 4 | AD | 1 | 73 | 97 |  | 0 | 2e-09 | 1197 | 1197 | -6.2785e-13
1 | 8 | 8 | 64 | 0 | DF | A1 | 98 | 8 | 87 | 12 |  | 100 | 0.02 | 14655896 | -2121320 | 0.011127 | Not Fully Discharged
1 | 8 | 8 | C8 | 0 | 0 | 4 | 7D | 8 | FE | 19 |  | 200 | 0.02 | 1149 | 1149 | -6.0267e-06
1 | 8 | 8 | C8 | 0 | F8 | 5E | 22 | 1 | 16 | 9C |  | 200 | 2e-09 | 16277026 | -500190 | 2.6236e-10
1 | 8 | 8 | 2C | 1 | D1 | 3A | E8 | 8 | E3 | 9D |  | 300 | 0.02 | 13712104 | -3065112 | 0.016077 | Not Fully Discharged
1 | 8 | 8 | 90 | 1 | FF | F7 | C6 | 8 | 7C | 16 |  | 400 | 0.02 | 16775110 | -2106 | 1.1046e-05
1 | 8 | 8 | 90 | 1 | ED | EA | C7 | 5 | 29 | 3D |  | 400 | 2e-05 | 15592135 | -1185081 | 6.216e-06
1 | 8 | 8 | F4 | 1 | BE | DD | CB | 5 | 84 | 93 |  | 500 | 2e-05 | 12508619 | -4268597 | 2.239e-05
1 | 8 | 8 | F4 | 1 | CF | 49 | B2 | 6 | BD | D2 |  | 500 | 0.0002 | 13584818 | -3192398 | 0.00016745
1 | 8 | 8 | 58 | 2 | C0 | F1 | BB | 6 | 64 | DF |  | 600 | 0.0002 | 12644795 | -4132421 | 0.00021675
1 | 8 | 8 | 58 | 2 | C0 | 0E | C3 | 7 | B7 | 2F |  | 600 | 0.002 | 12586691 | -4190525 | 0.002198
1 | 8 | 8 | 58 | 2 | F0 | DF | 0B | 8 | FF | D2 |  | 600 | 0.02 | 15785739 | -991477 | 0.0052005 | Not Fully Discharged
1 | 8 | 8 | BC | 2 | FD | AE | 9 | 8 | BB | E1 |  | 700 | 0.02 | 16625161 | -152055 | 0.00079756
1 | 8 | 8 | BC | 2 | FC | 40 | 9D | 7 | F4 | EC |  | 700 | 0.002 | 16531613 | -245603 | 0.00012882
1 | 8 | 8 | BC | 2 | E2 | 14 | 78 | 6 | 38 | 44 |  | 700 | 0.0002 | 14816376 | -1960840 | 0.00010285
1 | 8 | 8 | 20 | 3 | C0 | D0 | E4 | 6 | 3A | 9D |  | 800 | 0.0002 | 12636388 | -4140828 | 0.0002172
1 | 8 | 8 | 20 | 3 | C0 | 56 | 66 | 7 | 7A | 14 |  | 800 | 0.002 | 12605030 | -4172186 | 0.0021884
1 | 8 | 8 | 20 | 3 | D9 | 26 | 59 | 8 | 2D | 67 |  | 800 | 0.02 | 14231129 | -2546087 | 0.013355 | Not Fully Discharged
1 | 8 | 8 | 84 | 3 | FA | 94 | 27 | 8 | BE | 40 |  | 900 | 0.02 | 16421927 | -355289 | 0.0018636
1 | 8 | 8 | 84 | 3 | F9 | 25 | 28 | 7 | AB | D7 |  | 900 | 0.002 | 16327976 | -449240 | 0.00023564
1 | 8 | 8 | E8 | 3 | C3 | 26 | 86 | 7 | 22 | 3 |  | 1000 | 0.002 | 12789382 | -3987834 | 0.0020917
1 | 8 | 8 | E8 | 3 | D3 | A7 | AC | 8 | 29 | 8F |  | 1000 | 0.02 | 13871020 | -2906196 | 0.015244 | Not Fully Discharged
1 | 8 | 8 | 4C | 4 | F6 | FD | AE | 8 | AF | 14 |  | 1100 | 0.02 | 16186798 | -590418 | 0.0030969
1 | 8 | 8 | B0 | 4 | F6 | 8B | 82 | 8 | 46 | 32 |  | 1200 | 0.02 | 16157570 | -619646 | 0.0032502
1 | 8 | 8 | 14 | 5 | F4 | 68 | 41 | 8 | C2 | 68 |  | 1300 | 0.02 | 16017473 | -759743 | 0.003985
1 | 8 | 8 | 78 | 5 | F2 | 3D | A3 | 8 | 93 | 3C |  | 1400 | 0.02 | 15875491 | -901725 | 0.0047297
1 | 8 | 8 | DC | 5 | F0 | 9 | 38 | 8 | A1 | 5E |  | 1500 | 0.02 | 15731000 | -1046216 | 0.0054876
1 | 8 | 8 | 40 | 6 | ED | CC | 9A | 8 | 97 | F3 |  | 1600 | 0.02 | 15584410 | -1192806 | 0.0062565
1 | 8 | 8 | A4 | 6 | EB | C2 | 74 | 8 | AC | FC |  | 1700 | 0.02 | 15450740 | -1326476 | 0.0069576
1 | 8 | 8 | 8 | 7 | EB | 35 | 37 | 8 | 9 | 92 |  | 1800 | 0.02 | 15414583 | -1362633 | 0.0071473
1 | 8 | 8 | 6C | 7 | E8 | F3 | D7 | 8 | A8 | 0F |  | 1900 | 0.02 | 15266775 | -1510441 | 0.0079226
1 | 8 | 8 | D0 | 7 | E6 | AC | 38 | 8 | CC | 39 |  | 2000 | 0.02 | 15117368 | -1659848 | 0.0087062
1 | 8 | 8 | 34 | 8 | E4 | 63 | A5 | 8 | D6 | CB |  | 2100 | 0.02 | 14967717 | -1809499 | 0.0094912
1 | 8 | 8 | 98 | 8 | E2 | 20 | 4F | 8 | 70 | 9B |  | 2200 | 0.02 | 14819407 | -1957809 | 0.010269
1 | 8 | 8 | FC | 8 | E0 | 2B | BE | 8 | 4D | 55 |  | 2300 | 0.02 | 14691262 | -2085954 | 0.010941
1 | 8 | 8 | 60 | 9 | DF | 76 | EA | 8 | CE | 0F |  | 2400 | 0.02 | 14644970 | -2132246 | 0.011184
1 | 8 | 8 | C4 | 9 | DD | 1E | 68 | 8 | 37 | EF |  | 2500 | 0.02 | 14491240 | -2285976 | 0.01199
1 | 8 | 8 | 28 | 0A | DA | CC | A1 | 8 | 92 | 5E |  | 2600 | 0.02 | 14339233 | -2437983 | 0.012788
1 | 8 | 8 | 8C | 0A | D8 | 7F | 3C | 8 | 13 | B5 |  | 2700 | 0.02 | 14188348 | -2588868 | 0.013579
1 | 8 | 8 | F0 | 0A | D6 | 5D | 6C | 8 | 86 | AB |  | 2800 | 0.02 | 14048620 | -2728596 | 0.014312
1 | 8 | 8 | 54 | 0B | D5 | DA | 24 | 8 | 24 | 22 |  | 2900 | 0.02 | 14015012 | -2762204 | 0.014488
1 | 8 | 8 | B8 | 0B | D3 | 88 | A0 | 8 | F0 | D7 |  | 3000 | 0.02 | 13863072 | -2914144 | 0.015285
Problem 1 : there is time required to discharge to actual current whenever there is current range change. The larger the Gap between actual and Range might need long time to discharge, this issue will also cause inaccurate for the Autorange to determine correct Range to apply.